| Title Flashing Application and Bootloader Firmware |  |
| --- | --- |
| Document # | Version # |
| Effective Date | Prepared By |
| Revision Date | Approved By |

| Policy |  |
| --- | --- |
| Purpose |  |
| Scope |  |

| Materials | USB Cable Type-C Device Power | STM32CubeProgrammer firmwareupdateGUI.exe (v1.0.0.5) |
| --- | --- | --- |

Procedure

Required files

- Application firmware (pmfcAPP_XXiYYiZZ)

- Bootloader firmware (pmfcBL_XXiYYiZZ.hex)

Required software

- firmwareupdateGUI.exe (v1.0.0.5)

- STM32CubeProgrammer

Uploading Application Firmware

- Run firmwareupdateGUI.exe

- Click on Link->Device COM port

- Click on File->Firmware->Application

- Select the application firmware file

- Wait for the upgrade to complete. Progress bar is shown below

- Once upload completes, click OK on status message box. The new firmware version should appear in the lower left corner of the application.

Uploading Bootloader Firmware

- Click on File->Firmware->DFU Bootloader

- A warning BEEP and message will appear. Click OK to complete the process.

BOOTLOADER UPDATE CAN NOT BE INTERRUPTED!!!! OTHERWISE DEVICE NEEDS TO BE OPENED TO LAUNCH DFU FROM HARDWARE.

- Click OK on INFORMATION. Application software will close automatically.

- Opening Device Manager should show the device went to DFU Mode

- Run STM32CubeProgrammer

- Click on USB and Connect. Click on Download button then Browse. Select the bootloader file (pmfcBL_XXiYYiZZ.hex)

- Make sure Verify programming and Run after programming are checked. Click on Start Programming

- Connection from the device will be lost from STM32CuberProgrammer

- Click OK (there will be more messages but just click OK) and Close STM32CubeProgrammer

Verifying Application Firmware

- Use probeMFC or RHRII to check communications

Restore Points

- After the upgrade, device will create a restore point automatically based on what is currently saved in bulk flash data memory.

- To manually create a new restore point, click on File->Restore Point->Create Restore Point

- Once completed, software will prompt user that a restore point is created. Click OK to continue

Load Restore Points

- To load restore point, click on File->Restore Point->Load Restore Point

- Once complete, software will prompt user that the restore point was loaded and the image is saved in main internal data flash memory. Click OK to continue

NOTES: Restore points are created in manufacturing or in the field to create a “known good” flash page if configuration was changed by the user and want to revert to factory default.

Device with non-working COM port

- Unfortunately, these devices will need to be opened up to run DFU from hardware (by shorting DFU pins before applying power).

- Once device enters DFU mode, proceed with the step 5 to 7 of Uploading Bootloader Firmware

- Once completed, press and hold TARE button while applying power. Hold for 30 seconds and device should enter in boot mode (on display it says boot)

- Run the firmware update outlined under Uploading Application Firmware